**URGENT MEDICAL DEVICE NOTIFICATION – EIF-000561**

**Table of Contents**

[**URGENT MEDICAL DEVICE NOTIFICATION - US** 2](#_Toc165392222)

[**ACKNOWLEDGEMENT FORM - US** 6](#_Toc165392223)

[**URGENT – FIELD SAFETY NOTICE – EEA/TR/UK** 7](#_Toc165392224)

[**Annex 1: EIF-000561 – Impacted EEA/TR/UK POINTS OF CONTACT** 11](#_Toc165392225)

[**ACKNOWLEDGEMENT FORM – EEA/TR/UK** 12](#_Toc165392226)

[**APPENDIX 2: PRODUCT IN SCOPE OF EIF-000561** 13](#_Toc165392227)

Arrow International LLC

3015 Carrington Mill Blvd.

Morrisville, NC 27560

April-2024

## **URGENT MEDICAL DEVICE NOTIFICATION - US**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Type of Action** | **Correction** | | | |
| **Product Name** | **Arrow® FiberOptix™ Intra-Aortic Balloon Catheter Kit**  **Arrow® UltraFlex™ Intra-Aortic Balloon Catheter Kit** | | | |
| **Teleflex Reference** | **EIF-000561** | | | |
| **Product Code** | IAB-05830-LWS | IAB-05830-U | IAB-05840-LWS | IAB-05840-U |
| IAB-05850-LWS | IAB-06830-U | IAB-06840-U | IAB-06850-U |
| **Batch / Lot Number:** | **Refer to Appendix 2** | | | |
| **Unique Device Identifier (UDI)** |

Dear Customer,

**Details of affected devices**

Arrow International LLC, a subsidiary of Teleflex Incorporated (collectively “Teleflex”), has initiated a voluntary field correction for the above listed products; refer to Appendix 2 for Product Code, lot number and Unique Device Identifier (UDI) information.

**Description of the issue & immediate actions required**

Teleflex is initiating this voluntary correction for the above-mentioned products due to reports indicating an infrequent condition that, when not identified and corrected promptly, could result in serious health consequences. The issue may manifest as:

* failure of the intra-aortic balloon to completely inflate over its full length
* damaged or broken central lumen in the segment contained within the balloon
* helium loss or blood in the helium pathway.

Possible consequences of the above issue include a reduction or loss of the hemodynamic support normally provided by IABP therapy. This would require immediate replacement of the device resulting in delay or temporary interruption of IABP therapy. Therefore, this condition could result in an increased risk of hemodynamic or ischemic complications, including death. Between 05-April-2022 and 05-April-2024, a total of 337 reports of the above-mentioned conditions have been received. This includes reports of 31 injuries and 3 deaths. As reported, two of the three deaths were unrelated to the issue and the patient death occurred several days later. In the third event, the patient death may or may not have been related; the device was used in profound cardiogenic shock with repeated red-condition alarms.

**Actions to detect this issue:**

1. Prior to use, **ensure that a back up IAB catheter is available**, in the event that a replacement device is needed.
2. Prior to insertion, **users should inspect all in-scope IAB catheters** for signs of an over-twisted balloon wrap or bent balloon shaft. Over-twisting manifests as a progressive pitch or tightness within the helical wrap at either the distal or proximal end of the balloon. Examples of correct wrap versus over-twisted wrap are shown in images 1a & 1b below. Do not insert catheters suspected of manifesting an over-twisted wrap. **However, not all affected catheters will manifest with a visibly over-twisted balloon wrap.**

A close-up of a spiral object

Description automatically generated

*Image 1*

1. All in-scope IAB catheters must be inserted under fluoroscopic guidance. Once inserted and connected, **fluoroscopy must be used to assess the completeness of balloon inflation**. The fluoroscopic assessment must span the full length of the balloon (panning as necessary) and should include several cycles of inflation/deflation. Intermittent fluoroscopy should remain available for up to 3 minutes after one-to-one cyclic counterpulsation begins, or until full balloon inflation is confirmed.
2. As always, users should be vigilant in responding to pump alarms consistent with the Instructions For Use and Users Guide.

* In the event an Arrow AC3 or Autocat-2 pump is being used, a **high-pressure alarm** or **high baseline alarm** may sound and display on the pump console. These red-condition alarms require immediate attention per the Instructions For Use and Users Guide. The alarms provide users with a trouble shooting algorithm that includes incomplete IAB inflation.
* In the event a Getinge CardioSave pump is being used, an **IAB catheter restriction** alarm may sound.
* Any pump **alarm indicating helium leakage** (**helium loss** or **gas loss**) may signal the IAB catheter is not performing as expected.
* As always, real-time telephone support for Arrow IABPs and IAB catheters is available at the Patient Care Support contact details below.

1. As always, users should be vigilant in responding to bedside indicators that an IAB catheter is not performing as expected, including:

* Blood in the helium pathway
* Lower than expected diastolic pressure augmentation

**Immediate actions if this issue is identified:**

If an IAB catheter is suspected of having this issue, based upon any of the above features, the catheter should be removed and replaced using steps outlined in the Instructions for Use. A replacement balloon catheter may be inserted contralaterally or ipsilaterally as determined by the responsible physician after considering the patient’s individual femoral access options and relevant clinical features.

Every suspected incidence of this condition should be promptly reported to Teleflex using the contact information below as soon as practical. Your local sales representative can also assist with completing a report if needed.

Our records indicate you have received products that are subject to this corrective notice.

**Product is not being removed; you may continue to use the products in scope of this advisory notice per the additional instructions, warnings, and cautions contained herein.**

**Depending on your device location please adhere to the following Action list:**

|  |  |
| --- | --- |
| **Device location** | **Action List Number** |
| **Medical facilities (hospitals, medical staff, etc.)** | **1** |
| **Distributors** | **2** |

**Action list number 1 – Medical facilities**

1. Immediately provide this corrective notice to all persons who need to be aware within your organization (e.g., Coronary Care Unit, Interventional Cardiology Department, Cardiac Catheter Lab, etc.) or to any organization where the potentially affected product has been transferred.
2. We request that you immediately check your inventory for product within the scope of this Urgent Medical Device Notification.
3. If you do have affected product, mark the applicable checkbox on the Acknowledgement Form (Appendix 1) and return the form to Teleflex at the contact details provided, and place a copy of this notification with all affected product.
4. If you do not have affected product, mark the applicable checkbox on the Acknowledgement Form (Appendix 1) and return the form to Teleflex at the contact details provided.

**Action list number 2 – Distributors**

1. Immediately provide this Urgent Medical Device Notification to all customers who have received affected product. Request that your customer complete the Acknowledgement Form (Appendix 1) and return to you.
2. We request that you immediately check your inventory for affected product. Place a copy of this notification with all affected product and ensure that the notification is included with future distributions of affected product.
3. As a distributor, you are then required to confirm to Teleflex that you have completed the field activity outlined in actions 1 and 2 of this Action List Number 2. Upon completion of your actions, please forward the completed Acknowledgement Form to the e-mail address below.
4. If you have further distributed product outside of your country, please notify Teleflex Customer Service by return e-mail to the e-mail address below.

**Reporting Adverse Events**

Adverse events or quality issues experienced with the use of this product should be reported to Teleflex Customer Service using the contact details provided below and may also be reported to the FDA’s MedWatch Adverse Event Reporting program either by phone at 1-888-INFO-FDA (1-888-463-6332) or online at <https://www.fda.gov/safety/medwatch-fda-safety-information-and-adverse-event-reporting-program>.

**Transmission of this Corrective Notice**

This corrective notice should be given to all persons who need to be aware within your organization or to any organization where the potentially affected product have been transferred.

Please consider end users, clinicians, risk managers, supply chain/distribution centers, service departments, etc., in the circulation of this notice.

**Contact reference information**

As always, real-time telephone support for Arrow IABPs and IAB catheters is available at the Patient Care Support contact details below. Should you require any further information or support concerning this issue, please contact Teleflex Customer Service via email, phone, or fax. Customer Service hours of operation are 8am-7pm EST.

**24/7 Patient Care Support:**

**Telephone:** (800) 447-IABP (4227) USA/Canada

**Customer Service:**

**Contact:** Customer Service **Telephone:** 1-866-396-2111

**FAX:** 1-855-419-8507 **Email:** [Recalls@teleflex.com](mailto:Recalls@teleflex.com)

Teleflex is committed to providing high quality, safe and effective products. We regret any inconvenience this action may cause your operations. If you have any other questions, please contact your local Teleflex sales representative or Teleflex Customer Service.

***For and on behalf of Teleflex,***

*Padraig Hegarty*

***Padraig Hegarty VP, Global QA (Manufacturing)***

**Consignee No \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**Appendix 1**

### **ACKNOWLEDGEMENT FORM - US**

**MEDICAL DEVICE CORRECTIVE NOTICE BY TELEFLEX – IMMEDIATE ATTENTION REQUIRED**

**Ref. EIF-000561**

**RETURN COMPLETED FORM IMMEDIATELY TO:**

**E-mail:** [Recalls@teleflex.com](mailto:Recalls@teleflex.com)

|  |  |
| --- | --- |
| We confirm receipt of this corrective notice and completion of the required actions contained therein. We further confirm that our inventory does **NOT** include products affected by this corrective notice. | We confirm receipt of this corrective notice and completion of the required actions contained therein. We further confirm our inventory **DOES** include products affected by this corrective notice. |

**Complete this Acknowledgement Form and return the completed form immediately using the contact information above.**

|  |  |  |
| --- | --- | --- |
| **INSTITUTION NAME (E.G., NAME OF HOSPITAL, HEALTH CARE ORGANIZATION)** | | |
|  | | |
| **INSTITUTION ADDRESS** | | **PHONE/FAX** |
|  | |  |
| **FORM COMPLETED BY** | | **E-MAIL** |
| **PRINT NAME: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **SIGNATURE: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_** | |  |
| **DATE** |  |

Teleflex Medical

IDA Business & Technology Park,

Dublin Road, Athlone

Westmeath, Ireland

April-2024

## **URGENT – FIELD SAFETY NOTICE – EEA/TR/UK**

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Type of Action** | **Advisory Notice** | | | |
| **Product Name** | **Arrow® FiberOptix™ Intra-Aortic Balloon Catheter Kit**  **Arrow® UltraFlex™ Intra-Aortic Balloon Catheter Kit** | | | |
| **Teleflex Reference** | **EIF-000561** | | | |
| **Product Code** | IAB-05830-LWS | IAB-05830-U | IAB-05840-LWS | IAB-05840-U |
| IAB-05850-LWS | IAB-06830-U | IAB-06840-U | IAB-06850-U |
| **Batch / Lot Number:** | **Refer to Appendix 2** | | | |
| **Unique Device Identifier (UDI)** |

Dear Customer,

**Details of affected devices**

Arrow International LLC, a subsidiary of Teleflex Incorporated (collectively “Teleflex”), has initiated a voluntary Field Safety Corrective Action (“FSCA”) for the above listed products; refer to Appendix 2 for Product Code, lot number and Unique Device Identifier (UDI) information.

**Description of the** **issue & immediate actions required**

Teleflex is initiating this voluntary FSCA for the above-mentioned products due to reports indicating an infrequent condition that, **when not identified and corrected promptly, could result in serious health consequences**. The issue may manifest as:

* failure of the intra-aortic balloon to completely inflate over its full length
* damaged or broken central lumen in the segment contained within the balloon
* helium loss or blood in the helium pathway.

Possible consequences of the above issue include a reduction or loss of the hemodynamic support normally provided by IABP therapy. This would require immediate replacement of the device resulting in delay or temporary interruption of IABP therapy. **Therefore, this condition could result in an increased risk of hemodynamic or ischemic complications, including death.** Between 05-April-2022 and 05-April-2024, a total of 337 reports of the above-mentioned conditions have been received. This includes reports of 31 injuries and 3 deaths. As reported, two of the three deaths were unrelated to the issue and the patient death occurred several days later. In the third event, the patient death may or may not have been related; the device was used in profound cardiogenic shock with repeated red-condition alarms.

**Actions to detect this issue:**

1. Prior to use, **ensure that a back up IAB catheter is available**, in the event that a replacement device is needed.
2. Prior to insertion, **users should inspect all in-scope IAB catheters** for signs of an over-twisted balloon wrap or bent balloon shaft. Over-twisting manifests as a progressive pitch or tightness within the helical wrap at either the distal or proximal end of the balloon. Examples of correct wrap versus over-twisted wrap are shown in images 1a & 1b below. Do not insert catheters suspected of manifesting an over-twisted wrap. **However, not all affected catheters will manifest with a visibly over-twisted balloon wrap.**

A close-up of a spiral object

Description automatically generated

*Image 1*

1. All in-scope IAB catheters must be inserted under fluoroscopic guidance. Once inserted and connected, **fluoroscopy must be used to assess the completeness of balloon inflation**. The fluoroscopic assessment must span the full length of the balloon (panning as necessary) and should include several cycles of inflation/deflation. Intermittent fluoroscopy should remain available for up to 3 minutes after one-to-one cyclic counterpulsation begins, or until full balloon inflation is confirmed.
2. As always, users should be vigilant in responding to pump alarms consistent with the Instructions For Use and Users Guide.

* In the event an Arrow AC3 or Autocat-2 pump is being used, a **high-pressure alarm** or **high baseline alarm** may sound and display on the pump console. These red-condition alarms require immediate attention per the Instructions For Use and Users Guide. The alarms provide users with a trouble shooting algorithm that includes incomplete IAB inflation.
* In the event a Getinge CardioSave pump is being used, an **IAB catheter restriction** alarm may sound.
* Any pump **alarm indicating helium leakage** (**helium loss** or **gas loss**) may signal the IAB catheter is not performing as expected.
* As always, real-time telephone support for Arrow IABPs and IAB catheters is available at the Patient Care Support contact details below.

1. As always, users should be vigilant in responding to bedside indicators that an IAB catheter is not performing as expected, including:

* Blood in the helium pathway
* Lower than expected diastolic pressure augmentation

**Immediate actions if this issue is identified:**

If an IAB catheter is suspected of having this issue, based upon any of the above features, the catheter should be removed and replaced using steps outlined in the Instructions for Use. A replacement balloon catheter may be inserted contralaterally or ipsilaterally as determined by the responsible physician after considering the patient’s individual femoral access options and relevant clinical features.

Every suspected incidence of this condition should be promptly reported to Teleflex using the contact information below as soon as practical. Your local sales representative can also assist with completing a report if needed.

Our records indicate you have received products that are subject to this FSCA.

**Product is not being removed; you may continue to use the products in scope of this advisory notice per the additional instructions, warnings, and cautions contained herein.**

**Depending on your device location please adhere to the following Action list:**

|  |  |
| --- | --- |
| **Device location** | **Action List Number** |
| **Medical facilities (hospitals, medical staff, etc.)** | **1** |
| **Distributors** | **2** |

**Action list number 1 – Medical facilities**

1. Immediately provide this advisory notice to all persons who need to be aware within your organisation (e.g., Coronary Care Unit, Interventional Cardiology Department, Cardiac Catheter Lab, etc.) or to any organisation where the potentially affected product has been transferred.
2. We request that you immediately check your inventory for product within the scope of this Urgent Medical Device Notification.
3. If you have affected product, mark the applicable checkbox on the Acknowledgement Form (Appendix 1) and return the form to Teleflex Customer Service at the contact details provided, and place a copy of this notification with all affected product. If you would like to return affected product, this can also be initiated by contacting Teleflex Customer Service at the contact details provided below.
4. If you do not have affected product, mark the applicable checkbox on the Acknowledgement Form (Appendix 1) and return the form to Teleflex at the contact details provided.

**Action list number 2 – Distributors**

1. Update the contact details on pages ten and twelve from the Teleflex customer service contact details to your contact details.
2. Immediately provide this field safety notice to all customers who have received affected product. Request that your customer complete the Acknowledgement Form (Appendix 1) and return to you.
3. We request that you immediately check your inventory for affected product in scope of this FSCA, refer to Appendix 2 for the list of affected codes and lots. Place a copy of this notification with all affected product and ensure that the notification is included with future distributions of affected product.
4. As a distributor, you are then required to confirm to Teleflex that you have completed the field activity outlined in actions 1 and 2 of this Action List Number 2. Upon completion of your actions, please forward the completed Acknowledgement Form to Customer Service.
5. Please be aware that all European Economic Area/Switzerland/United Kingdom (EEA/CH/UK) and Turkey Member State Competent Authorities in which Teleflex distributes directly will be notified by Teleflex.
6. If you have further distributed product outside of your country, please notify Teleflex Customer Service by return e-mail to the e-mail address below.
7. If you are a distributor and/or have a reporting responsibility within or outside the EEA/CH/TR/UK region, please notify your local competent authority of this action. Please forward the notification and all communication with your local competent authority to Teleflex.

**Reporting Adverse Events**

Adverse events or quality problems experienced with the use of this product should be reported to Teleflex Customer Service using the contact details provided below.

**Transmission of this Advisory Notice**

This advisory notice should be given to all persons who need to be aware within your organisation or to any organisation where the potentially affected product have been transferred.

Please consider end users, clinicians, risk managers, supply chain/distribution centers, service departments, etc., in the circulation of this notice. Maintain awareness of this notice until all required actions have been completed in your organisation.

**Contact reference information**

As always, real-time telephone support for Arrow IABPs and IAB catheters is available at the Patient Care Support contact details below. Should you require any further information or support concerning this issue, please contact Teleflex Customer Service via email, phone, or fax. Customer Service hours of operation are 8am-7pm EST.

**24/7 Patient Care Support:**

**Telephone:** (978) 250-5177 International

**Contact reference person:**

Should you require any further information or support concerning this issue, please referto**:**

**Annex 1 (page 11) for the applicable EEA/TR/UK point of contact name, email address and telephone number.**

Teleflex is committed to providing high quality, safe and effective products. We regret any inconvenience this action may cause your operations. If you have any other questions, please contact your local Teleflex sales representative or Teleflex Customer Service.

The undersign confirms this notice has been notified to the appropriate Regulatory Authority.

***For and on behalf of Teleflex,***

*Padraig Hegarty*

***Padraig Hegarty VP, Global QA (Manufacturing)***

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| **Annex 1: EIF-000561 – Impacted EEA/TR/UK POINTS OF CONTACT** | | | | |
| ***EMEA Country*** | ***Point of Contact*** | ***e-mail*** | ***Phone*** |
| ***Austria*** | Customer Service | [qualityAT@teleflex.com](mailto:qualityAT@teleflex.com) | +43 0 1 402 4772 DW 94 |
| ***Belgium*** | Sales Assistant | [productcomplaints.belgium@teleflex.com](mailto:productcomplaints.belgium@teleflex.com) | +32 (02) 333 24 60 |
| ***Bulgaria*** | Shane Kenny | [Recalls.Intl@teleflex.com](mailto:Recalls.Intl@teleflex.com) | +353 (0)86 3479154 |
| ***Cyprus*** | Shane Kenny | [Recalls.Intl@teleflex.com](mailto:Recalls.Intl@teleflex.com) | +353 (0)86 3479154 |
| ***Czech*** | Maria Hejcmanova | [maria.hejcmanova@teleflex.com](mailto:maria.hejcmanova@teleflex.com) | +420 601 203 708 |
| ***Germany*** | Customer Service | [recalls.de@teleflex.com](mailto:recalls.de@teleflex.com) | 0711 2090-8000 |
| ***Spain*** | Gloria Blanco | [gloria.blanco@teleflex.com](mailto:gloria.blanco@teleflex.com) | +34 649392681;  +34 918300462 |
| ***France*** | Carine Fournier | [productcomplaints.france@teleflex.com](mailto:productcomplaints.france@teleflex.com) | 05 62 18 79 41 |
| ***Great Britain*** | Cliodhna Coffey | [Recalls.uk@teleflex.com](mailto:Recalls.uk@teleflex.com) | +44 (0) 1494 532761 |
| ***Greece*** | Maria Halatsi | [maria.halatsi@teleflex.com](mailto:maria.halatsi@teleflex.com) | + 302106777717 |
| ***Hungary*** | Shane Kenny | [Recalls.Intl@teleflex.com](mailto:Recalls.Intl@teleflex.com) | +353 (0)86 3479154 |
| ***Ireland*** | Cliodhna Coffey | [Recalls.ie@teleflex.com](mailto:Recalls.ie@teleflex.com) | +353 906460940 |
| ***Italy*** | Giovanni Cordone | [giovanni.cordone@teleflex.com](mailto:giovanni.cordone@teleflex.com) | 3625890252 |
| ***Luxembourg*** | Sales Assistant | [productcomplaints.belgium@teleflex.com](mailto:productcomplaints.belgium@teleflex.com) | +32 (02) 333 24 60 |
| ***Netherlands*** | Sales Assistant | [productcomplaints.netherlands@teleflex.com](mailto:productcomplaints.netherlands@teleflex.com) | +31 (0) 088 00 215 00 |
| ***Poland*** | Katarzyne Demczuk | [orders.pl@teleflex.com](mailto:orders.pl@teleflex.com) | +48 22 462 40 32 |
| ***Portugal*** | Gloria Blanco | [gloria.blanco@teleflex.com](mailto:gloria.blanco@teleflex.com) | +34 918300432;  +34 649392681 |
| ***Romania*** | Shane Kenny | [Recalls.Intl@teleflex.com](mailto:Recalls.Intl@teleflex.com) | +353 (0)86 3479154 |
| ***Sweden*** | Shane Kenny | [Recalls.Intl@teleflex.com](mailto:Recalls.Intl@teleflex.com) | +353 (0)86 3479154 |
| ***Slovenia*** | Shane Kenny | [Recalls.Intl@teleflex.com](mailto:Recalls.Intl@teleflex.com) | +353 (0)86 3479154 |
| ***Slovakia*** | Erika Mikolasikova | [erika.mikolasikova@teleflex.com](mailto:erika.mikolasikova@teleflex.com) | +421337725428 |
| ***Turkey*** | Shane Kenny | [Recalls.Intl@teleflex.com](mailto:Recalls.Intl@teleflex.com) | +353 (0)86 3479154 |
| Note: For a translated copy of the Field Safety Notice, please contact the applicable EEA/TR/UK point of contact | | | | |

**Customer No \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**

**Appendix 1**

**FIELD SAFETY CORRECTIVE ACTION**

### **ACKNOWLEDGEMENT FORM – EEA/TR/UK**

**PRODUCT FIELD ACTION NOTICE BY TELEFLEX – IMMEDIATE ATTENTION REQUIRED**

**Ref. EIF-000561**

**RETURN COMPLETED FORM IMMEDIATELY TO:**

**E-Mail:** ***Refer to Annex 1 for the applicable point of contact email address.***

|  |  |
| --- | --- |
| We confirm receipt of this advisory notice and completion of the required actions contained therein. We further confirm that our inventory does **NOT** include products affected by this advisory notice. | We confirm receipt of this advisory notice and completion of the required actions contained therein. We further confirm our inventory **DOES** include products affected by this advisory notice. |

**Complete this Acknowledgement Form and return the completed form immediately using the contact information above.**

|  |  |  |
| --- | --- | --- |
| **INSTITUTION NAME (E.G., NAME OF HOSPITAL, HEALTH CARE ORGANISATION)** | | |
|  | | |
| **INSTITUTION ADDRESS** | | **PHONE/FAX** |
|  | |  |
| **FORM COMPLETED BY** | | **Stamp** |
| **PRINT NAME: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_**  **SIGNATURE: \_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_\_** | |  |
| **DATE** |  |

## **APPENDIX 2: PRODUCT IN SCOPE OF EIF-000561**

| **Appendix 2: EIF-000561 - Product code, Lot Number, Unique Device Identifier (UDI) Information** | | |
| --- | --- | --- |
| **Product Code** | **Lot Number** | **Unique Device Identifier (UDI)** |
| IAB-05830-LWS | 18F23A0038 | (01)10801902161892(17)241231(10)18F23A0038 |
| IAB-05830-LWS | 18F23B0023 | (01)10801902161892(17)250131(10)18F23B0023 |
| IAB-05830-LWS | 18F23C0026 | (01)10801902172843(17)250228(10)18F23C0026 |
| IAB-05830-LWS | 18F23D0024 | (01)10801902161892(17)250430(10)18F23D0024 |
| IAB-05830-LWS | 18F23F0049 | (01)10801902172843(17)250531(10)18F23F0049 |
| IAB-05830-LWS | 18F23H0060 | (01)10801902161892(17)250731(10)18F23H0060 |
| IAB-05830-LWS | 18F23J0001 | (01)10801902172843(17)250831(10)18F23J0001 |
| IAB-05830-LWS | 18F23J0075 | (01)10801902161892(17)250930(10)18F23J0075 |
| IAB-05830-LWS | 18F23K0013 | (01)10801902161892(17)250930(10)18F23K0013 |
| IAB-05830-LWS | 18F23L0019 | (01)10801902161892(17)251031(10)18F23L0019 |
| IAB-05830-LWS | 18F23L0041 | (01)10801902161892(17)251031(10)18F23L0041 |
| IAB-05830-LWS | 18F23L0054 | (01)10801902172843(17)251031(10)18F23L0054 |
| IAB-05830-LWS | 18F23M0056 | (01)10801902172843(17)251130(10)18F23M0056 |
| IAB-05830-LWS | 18F24A0015 | (01)10801902172843(17)251231(10)18F24A0015 |
| IAB-05830-LWS | 18F24A0040 | (01)10801902172843(17)260131(10)18F24A0040 |
| IAB-05830-LWS | 18F24B0008 | (01)10801902161892(17)260131(10)18F24B0008 |
| IAB-05830-LWS | 18F24B0013 | (01)10801902172843(17)260131(10)18F24B0013 |
| IAB-05830-LWS | 18F24C0058 | (01)10801902161892(17)260228(10)18F24C0058 |
| IAB-05830-LWS | 18F24C0061 | (01)10801902172843(17)260228(10)18F24C0061 |
| IAB-05830-LWS | 18F24C0088 | (01)10801902172843(17)46112(10)18F24C0088 |
| IAB-05830-LWS | 18F24C0089 | (01)10801902172843(17)46112(10)18F24C0089 |
| IAB-05830-LWS | 18F24C0101 | (01)10801902161892(17)260228(10)18F24C0101 |
| IAB-05830-LWS | 18F24C0109 | (01)10801902172843(17)260228(10)18F24C0109 |
| IAB-05830-LWS | 18F24C0110 | (01)10801902172843(17)260228(10)18F24C0110 |
| IAB-05830-LWS | 18F24C0172 | (01)10801902172843(17)260331(10)18F24C0172 |
| IAB-05830-LWS | 18F24C0173 | (01)10801902172843(17)260331(10)18F24C0173 |
| IAB-05830-LWS | 18F24C0174 | (01)10801902172843(17)46112(10)18F24C0174 |
| IAB-05830-LWS | 18F24C0175 | (01)10801902172843(17)46112(10)18F24C0175 |
| IAB-05830-LWS | 18F24C0176 | (01)10801902172843(17)46112(10)18F24C0176 |
| IAB-05830-LWS | 18F24C0185 | (01)10801902172843(17)46112(10)18F24C0185 |
| IAB-05830-LWS | 18F24D0002 | (01)10801902161892(17)260331(10)18F24D0002 |
| IAB-05830-U | 18F22F0023 | (01)10801902172850(17)240531(10)18F22F0023 |
| IAB-05830-U | 18F22G0005 | (01)10801902172850(17)240630(10)18F22G0005 |
| IAB-05830-U | 18F22G0010 | (01)10801902172850(17)240630(10)18F22G0010 |
| IAB-05830-U | 18F22H0033 | (01)10801902172850(17)240731(10)18F22H0033 |
| IAB-05830-U | 18F23E0032 | (01)10801902182019(17)250430(10)18F23E0032 |
| IAB-05830-U | 18F23J0002 | (01)10801902182019(17)250831(10)18F23J0002 |
| IAB-05830-U | 18F23J0031 | (01)10801902172850(17)250831(10)18F23J0031 |
| IAB-05830-U | 18F23J0082 | (01)10801902182019(17)250831(10)18F23J0082 |
| IAB-05830-U | 18F23K0025 | (01)10801902182019(17)250930(10)18F23K0025 |
| IAB-05830-U | 18F23K0036 | (01)10801902172850(17)250930(10)18F23K0036 |
| IAB-05830-U | 18F23K0042 | (01)10801902172850(17)250930(10)18F23K0042 |
| IAB-05830-U | 18F23M0001 | (01)10801902172850(17)251130(10)18F23M0001 |
| IAB-05830-U | 18F24C0131 | (01)10801902182019(17)260331(10)18F24C0131 |
| IAB-05840-LWS | 18F22E0008 | (01)10801902172867(17)240430(10)18F22E0008 |
| IAB-05840-LWS | 18F22E0029 | (01)10801902172867(17)240430(10)18F22E0029 |
| IAB-05840-LWS | 18F22F0006 | (01)10801902172867(17)240531(10)18F22F0006 |
| IAB-05840-LWS | 18F22F0018 | (01)10801902172867(17)240531(10)18F22F0018 |
| IAB-05840-LWS | 18F22F0039 | (01)10801902172867(17)240531(10)18F22F0039 |
| IAB-05840-LWS | 18F22F0072 | (01)10801902167733(17)240531(10)18F22F0072 |
| IAB-05840-LWS | 18F22G0002 | (01)10801902172867(17)240630(10)18F22G0002 |
| IAB-05840-LWS | 18F22G0003 | (01)10801902172867(17)240630(10)18F22G0003 |
| IAB-05840-LWS | 18F22G0019 | (01)10801902172867(17)240630(10)18F22G0019 |
| IAB-05840-LWS | 18F22G0021 | (01)10801902167733(17)240630(10)18F22G0021 |
| IAB-05840-LWS | 18F22G0031 | (01)10801902167733(17)240630(10)18F22G0031 |
| IAB-05840-LWS | 18F22G0046 | (01)10801902167733(17)240630(10)18F22G0046 |
| IAB-05840-LWS | 18F22G0059 | (01)10801902172867(17)240630(10)18F22G0059 |
| IAB-05840-LWS | 18F22H0009 | (01)10801902167733(17)240731(10)18F22H0009 |
| IAB-05840-LWS | 18F22H0023 | (01)10801902167733(17)240731(10)18F22H0023 |
| IAB-05840-LWS | 18F22H0035 | (01)10801902172867(17)240731(10)18F22H0035 |
| IAB-05840-LWS | 18F22H0058 | (01)10801902167733(17)240731(10)18F22H0058 |
| IAB-05840-LWS | 18F22J0025 | (01)10801902172867(17)240831(10)18F22J0025 |
| IAB-05840-LWS | 18F22K0015 | (01)10801902167733(17)240930(10)18F22K0015 |
| IAB-05840-LWS | 18F22K0016 | (01)10801902172867(17)240930(10)18F22K0016 |
| IAB-05840-LWS | 18F22M0017 | (01)10801902167733(17)241130(10)18F22M0017 |
| IAB-05840-LWS | 18F22M0018 | (01)10801902167733(17)241130(10)18F22M0018 |
| IAB-05840-LWS | 18F23A0004 | (01)10801902167733(17)241231(10)18F23A0004 |
| IAB-05840-LWS | 18F23A0005 | (01)10801902167733(17)241231(10)18F23A0005 |
| IAB-05840-LWS | 18F23A0018 | (01)10801902167733(17)241231(10)18F23A0018 |
| IAB-05840-LWS | 18F23A0026 | (01)10801902167733(17)241231(10)18F23A0026 |
| IAB-05840-LWS | 18F23B0008 | (01)10801902167733(17)250131(10)18F23B0008 |
| IAB-05840-LWS | 18F23B0016 | (01)10801902167733(17)250131(10)18F23B0016 |
| IAB-05840-LWS | 18F23E0020 | (01)10801902172867(17)250430(10)18F23E0020 |
| IAB-05840-LWS | 18F23E0054 | (01)10801902167733(17)250531(10)18F23E0054 |
| IAB-05840-LWS | 18F23F0002 | (01)10801902172867(17)250531(10)18F23F0002 |
| IAB-05840-LWS | 18F23F0026 | (01)10801902172867(17)250531(10)18F23F0026 |
| IAB-05840-LWS | 18F23G0008 | (01)10801902172867(17)250630(10)18F23G0008 |
| IAB-05840-LWS | 18F23G0017 | (01)10801902167733(17)250630(10)18F23G0017 |
| IAB-05840-LWS | 18F23G0057 | (01)10801902172867(17)250731(10)18F23G0057 |
| IAB-05840-LWS | 18F23H0013 | (01)10801902172867(17)250731(10)18F23H0013 |
| IAB-05840-LWS | 18F23H0029 | (01)10801902167733(17)250731(10)18F23H0029 |
| IAB-05840-LWS | 18F23H0055 | (01)10801902172867(17)250731(10)18F23H0055 |
| IAB-05840-LWS | 18F23H0064 | (01)10801902167733(17)250731(10)18F23H0064 |
| IAB-05840-LWS | 18F23J0003 | (01)10801902172867(17)250831(10)18F23J0003 |
| IAB-05840-LWS | 18F23J0037 | (01)10801902172867(17)250831(10)18F23J0037 |
| IAB-05840-LWS | 18F23J0072 | (01)10801902172867(17)250831(10)18F23J0072 |
| IAB-05840-LWS | 18F23J0073 | (01)10801902172867(17)250930(10)18F23J0073 |
| IAB-05840-LWS | 18F23K0054 | (01)10801902172867(17)251031(10)18F23K0054 |
| IAB-05840-LWS | 18F23L0009 | (01)10801902167733(17)251031(10)18F23L0009 |
| IAB-05840-LWS | 18F23L0028 | (01)10801902167733(17)251031(10)18F23L0028 |
| IAB-05840-LWS | 18F23L0055 | (01)10801902172867(17)251130(10)18F23L0055 |
| IAB-05840-LWS | 18F23M0043 | (01)10801902172867(17)251130(10)18F23M0043 |
| IAB-05840-LWS | 18F23M0068 | (01)10801902167733(17)251231(10)18F23M0068 |
| IAB-05840-LWS | 18F24A0025 | (01)10801902172867(17)251231(10)18F24A0025 |
| IAB-05840-LWS | 18F24A0030 | (01)10801902167733(17)251231(10)18F24A0030 |
| IAB-05840-LWS | 18F24A0031 | (01)10801902172867(17)251231(10)18F24A0031 |
| IAB-05840-LWS | 18F24A0039 | (01)10801902167733(17)251231(10)18F24A0039 |
| IAB-05840-LWS | 18F24B0005 | (01)10801902172867(17)260131(10)18F24B0005 |
| IAB-05840-LWS | 18F24B0015 | (01)10801902167733(17)260131(10)18F24B0015 |
| IAB-05840-LWS | 18F24B0031 | (01)10801902167733(17)260131(10)18F24B0031 |
| IAB-05840-LWS | 18F24B0048 | (01)10801902172867(17)260131(10)18F24B0048 |
| IAB-05840-LWS | 18F24B0052 | (01)10801902172867(17)260131(10)18F24B0052 |
| IAB-05840-LWS | 18F24B0078 | (01)10801902172867(17)260228(10)18F24B0078 |
| IAB-05840-LWS | 18F24C0008 | (01)10801902167733(17)260228(10)18F24C0008 |
| IAB-05840-LWS | 18F24C0036 | (01)10801902172867(17)260228(10)18F24C0036 |
| IAB-05840-LWS | 18F24C0064 | (01)10801902172867(17)260228(10)18F24C0064 |
| IAB-05840-LWS | 18F24C0065 | (01)10801902172867(17)260228(10)18F24C0065 |
| IAB-05840-LWS | 18F24C0067 | (01)10801902172867(17)260228(10)18F24C0067 |
| IAB-05840-LWS | 18F24C0068 | (01)10801902172867(17)260228(10)18F24C0068 |
| IAB-05840-LWS | 18F24C0069 | (01)10801902167733(17)260228(10)18F24C0069 |
| IAB-05840-LWS | 18F24C0071 | (01)10801902167733(17)260228(10)18F24C0071 |
| IAB-05840-LWS | 18F24C0072 | (01)10801902167733(17)260228(10)18F24C0072 |
| IAB-05840-LWS | 18F24C0075 | (01)10801902167733(17)260228(10)18F24C0075 |
| IAB-05840-LWS | 18F24C0076 | (01)10801902167733(17)260228(10)18F24C0076 |
| IAB-05840-LWS | 18F24C0077 | (01)10801902167733(17)260228(10)18F24C0077 |
| IAB-05840-LWS | 18F24C0078 | (01)10801902167733(17)260228(10)18F24C0078 |
| IAB-05840-LWS | 18F24C0079 | (01)10801902167733(17)260228(10)18F24C0079 |
| IAB-05840-LWS | 18F24C0080 | (01)10801902167733(17)260228(10)18F24C0080 |
| IAB-05840-LWS | 18F24C0081 | (01)10801902167733(17)260228(10)18F24C0081 |
| IAB-05840-LWS | 18F24C0083 | (01)10801902167733(17)260331(10)18F24C0083 |
| IAB-05840-LWS | 18F24C0105 | (01)10801902172867(17)260228(10)18F24C0105 |
| IAB-05840-LWS | 18F24C0186 | (01)10801902167733(17)260331(10)18F24C0186 |
| IAB-05840-LWS | 18F24C0187 | (01)10801902167733(17)260331(10)18F24C0187 |
| IAB-05840-LWS | 18F24C0188 | (01)10801902167733(17)260331(10)18F24C0188 |
| IAB-05840-LWS | 18F24C0193 | (01)10801902167733(17)260331(10)18F24C0193 |
| IAB-05840-LWS | 18F24D0020 | (01)10801902167733(17)260331(10)18F24D0020 |
| IAB-05840-LWS | 18F24D0021 | (01)10801902167733(17)260331(10)18F24D0021 |
| IAB-05840-LWS | 18F24D0022 | (01)10801902167733(17)260331(10)18F24D0022 |
| IAB-05840-LWS | 18F24D0031 | (01)10801902167733(17)260331(10)18F24D0031 |
| IAB-05840-LWS | 18F24D0032 | (01)10801902167733(17)260331(10)18F24D0032 |
| IAB-05840-LWS | 18F24D0037 | (01)10801902167733(17)260331(10)18F24D0037 |
| IAB-05840-LWS | 18F24D0038 | (01)10801902167733(17)260331(10)18F24D0038 |
| IAB-05840-LWS | 18F24D0039 | (01)10801902167733(17)260331(10)18F24D0039 |
| IAB-05840-LWS | 18F24D0042 | (01)10801902167733(17)260331(10)18F24D0042 |
| IAB-05840-LWS | 18F24D0044 | (01)10801902167733(17)260331(10)18F24D0044 |
| IAB-05840-LWS | 18F24D0045 | (01)10801902167733(17)260331(10)18F24D0045 |
| IAB-05840-LWS | 18F24D0049 | (01)10801902167733(17)260331(10)18F24D0049 |
| IAB-05840-LWS | 18F24D0050 | (01)10801902167733(17)260331(10)18F24D0050 |
| IAB-05840-LWS | 18F24D0051 | (01)10801902167733(17)260331(10)18F24D0051 |
| IAB-05840-LWS | 18F24D0054 | (01)10801902167733(17)260331(10)18F24D0054 |
| IAB-05840-LWS | 18F24D0055 | (01)10801902167733(17)260331(10)18F24D0055 |
| IAB-05840-LWS | 18F24D0072 | (01)10801902167733(17)260331(10)18F24D0072 |
| IAB-05840-LWS | 18F24D0087 | (01)10801902167733(17)260331(10)18F24D0087 |
| IAB-05840-LWS | 18F24D0088 | (01)10801902167733(17)260331(10)18F24D0088 |
| IAB-05840-LWS | 18F24D0089 | (01)10801902167733(17)260331(10)18F24D0089 |
| IAB-05840-LWS | 18F24D0090 | (01)10801902167733(17)260331(10)18F24D0090 |
| IAB-05840-LWS | 18F24D0091 | (01)10801902167733(17)260331(10)18F24D0091 |
| IAB-05840-LWS | 18F24D0092 | (01)10801902167733(17)260331(10)18F24D0092 |
| IAB-05840-U | 18F22E0012 | (01)10801902172874(17)240430(10)18F22E0012 |
| IAB-05840-U | 18F22G0011 | (01)10801902172874(17)240630(10)18F22G0011 |
| IAB-05840-U | 18F22G0016 | (01)10801902172874(17)240630(10)18F22G0016 |
| IAB-05840-U | 18F22G0051 | (01)10801902161922(17)240630(10)18F22G0051 |
| IAB-05840-U | 18F23A0002 | (01)10801902161922(17)241231(10)18F23A0002 |
| IAB-05840-U | 18F23A0021 | (01)10801902161922(17)241231(10)18F23A0021 |
| IAB-05840-U | 18F23A0060 | (01)10801902161922(17)241231(10)18F23A0060 |
| IAB-05840-U | 18F23G0068 | (01)10801902161922(17)250731(10)18F23G0068 |
| IAB-05840-U | 18F23K0008 | (01)10801902172874(17)251130(10)18F23K0008 |
| IAB-05840-U | 18F23K0012 | (01)10801902161922(17)250930(10)18F23K0012 |
| IAB-05840-U | 18F23K0039 | (01)10801902172874(17)250930(10)18F23K0039 |
| IAB-05840-U | 18F23K0053 | (01)10801902161922(17)251031(10)18F23K0053 |
| IAB-05840-U | 18F23K0063 | (01)10801902161922(17)251031(10)18F23K0063 |
| IAB-05840-U | 18F23L0010 | (01)10801902172874(17)251031(10)18F23L0010 |
| IAB-05840-U | 18F23L0031 | (01)10801902172874(17)251031(10)18F23L0031 |
| IAB-05840-U | 18F23L0039 | (01)10801902172874(17)251031(10)18F23L0039 |
| IAB-05840-U | 18F24B0051 | (01)10801902161922(17)260131(10)18F24B0051 |
| IAB-05850-LWS | 18F22J0037 | (01)10801902167726(17)240831(10)18F22J0037(13)220927 |
| IAB-05850-LWS | 18F22J0040 | (01)10801902167726(17)240831(10)18F22J0040(13)220929 |
| IAB-05850-LWS | 18F22K0044 | (01)10801902167726(17)240930(10)18F22K0044(13)221021 |
| IAB-05850-LWS | 18F22K0045 | (01)10801902167726(17)240930(10)18F22K0045(13)221021 |
| IAB-05850-LWS | 18F22K0056 | (01)10801902167726(17)241031(10)18F22K0056(13)221028 |
| IAB-05850-LWS | 18F22L0012 | (01)10801902167726(17)241031(10)18F22L0012(13)221104 |
| IAB-05850-LWS | 18F23K0006 | (01)10801902167726(17)250930(10)18F23K0006(13)231031 |
| IAB-05850-LWS | 18F23K0038 | (01)10801902172881(17)250930(10)18F23K0038(13)231028 |
| IAB-05850-LWS | 18F24B0049 | (01)10801902167726(17)260131(10)18F24B0049(13)240227 |
| IAB-05850-LWS | 18F24B0075 | (01)10801902167726(17)260228(10)18F24B0075(13)240301 |
| IAB-05850-LWS | 18F24C0031 | (01)10801902172881(17)260228(10)18F24C0031(13)240316 |
| IAB-05850-LWS | 18F24C0032 | (01)10801902172881(17)260228(10)18F24C0032(13)240331 |
| IAB-05850-LWS | 18F24C0033 | (01)10801902172881(17)260228(10)18F24C0033(13)240315 |
| IAB-05850-LWS | 18F24C0034 | (01)10801902172881(17)260228(10)18F24C0034(13)240315 |
| IAB-06830-U | 18F22E0062 | (01)10801902172898(17)240430(10)18F22E0062 |
| IAB-06830-U | 18F22F0007 | (01)10801902172898(17)240531(10)18F22F0007 |
| IAB-06830-U | 18F22F0017 | (01)10801902172898(17)240531(10)18F22F0017 |
| IAB-06830-U | 18F22F0020 | (01)10801902172898(17)240531(10)18F22F0020 |
| IAB-06830-U | 18F22F0034 | (01)10801902172898(17)240531(10)18F22F0034 |
| IAB-06830-U | 18F22F0045 | (01)10801902172898(17)240531(10)18F22F0045 |
| IAB-06830-U | 18F22F0057 | (01)10801902172898(17)240531(10)18F22F0057 |
| IAB-06830-U | 18F22G0001 | (01)10801902172898(17)240630(10)18F22G0001 |
| IAB-06830-U | 18F22G0004 | (01)10801902172898(17)240630(10)18F22G0004 |
| IAB-06830-U | 18F22G0007 | (01)10801902172898(17)240630(10)18F22G0007 |
| IAB-06830-U | 18F22G0014 | (01)10801902172898(17)240630(10)18F22G0014 |
| IAB-06830-U | 18F22H0002 | (01)10801902172898(17)240731(10)18F22H0002 |
| IAB-06830-U | 18F22H0008 | (01)10801902161939(17)240731(10)18F22H0008 |
| IAB-06830-U | 18F22H0024 | (01)10801902172898(17)240731(10)18F22H0024 |
| IAB-06830-U | 18F22H0029 | (01)10801902172898(17)240731(10)18F22H0029 |
| IAB-06830-U | 18F22H0034 | (01)10801902172898(17)240731(10)18F22H0034 |
| IAB-06830-U | 18F22J0009 | (01)10801902172898(17)240831(10)18F22J0009 |
| IAB-06830-U | 18F22M0019 | (01)10801902161939(17)241130(10)18F22M0019 |
| IAB-06830-U | 18F23A0027 | (01)10801902161939(17)241231(10)18F23A0027 |
| IAB-06830-U | 18F23A0039 | (01)10801902161939(17)241231(10)18F23A0039 |
| IAB-06830-U | 18F23B0006 | (01)10801902161939(17)250131(10)18F23B0006 |
| IAB-06830-U | 18F23B0020 | (01)10801902161939(17)250131(10)18F23B0020 |
| IAB-06830-U | 18F23B0032 | (01)10801902161939(17)250131(10)18F23B0032 |
| IAB-06830-U | 18F23D0029 | (01)10801902172898(17)250430(10)18F23D0029 |
| IAB-06830-U | 18F23E0033 | (01)10801902172898(17)250531(10)18F23E0033 |
| IAB-06830-U | 18F23E0037 | (01)10801902172898(17)250430(10)18F23E0037 |
| IAB-06830-U | 18F23E0038 | (01)10801902172898(17)250430(10)18F23E0038 |
| IAB-06830-U | 18F23E0057 | (01)10801902172898(17)250531(10)18F23E0057 |
| IAB-06830-U | 18F23F0023 | (01)10801902172898(17)250531(10)18F23F0023 |
| IAB-06830-U | 18F23F0045 | (01)10801902172898(17)250531(10)18F23F0045 |
| IAB-06830-U | 18F23G0025 | (01)10801902172898(17)250630(10)18F23G0025 |
| IAB-06830-U | 18F23G0026 | (01)10801902172898(17)250630(10)18F23G0026 |
| IAB-06830-U | 18F23G0051 | (01)10801902172898(17)250630(10)18F23G0051 |
| IAB-06830-U | 18F23G0066 | (01)10801902161939(17)250630(10)18F23G0066 |
| IAB-06830-U | 18F23G0067 | (01)10801902172898(17)250731(10)18F23G0067 |
| IAB-06830-U | 18F23H0009 | (01)10801902172898(17)250731(10)18F23H0009 |
| IAB-06830-U | 18F23H0011 | (01)10801902172898(17)250831(10)18F23H0011 |
| IAB-06830-U | 18F23H0025 | (01)10801902172898(17)250731(10)18F23H0025 |
| IAB-06830-U | 18F23H0028 | (01)10801902172898(17)250731(10)18F23H0028 |
| IAB-06830-U | 18F23H0059 | (01)10801902161939(17)250731(10)18F23H0059 |
| IAB-06830-U | 18F23H0069 | (01)10801902172898(17)250731(10)18F23H0069 |
| IAB-06830-U | 18F23H0070 | (01)10801902161939(17)250731(10)18F23H0070 |
| IAB-06830-U | 18F23H0074 | (01)10801902161939(17)250731(10)18F23H0074 |
| IAB-06830-U | 18F23H0080 | (01)10801902172898(17)250731(10)18F23H0080 |
| IAB-06830-U | 18F23J0004 | (01)10801902172898(17)250831(10)18F23J0004 |
| IAB-06830-U | 18F23J0011 | (01)10801902172898(17)250831(10)18F23J0011 |
| IAB-06830-U | 18F23J0015 | (01)10801902172898(17)250831(10)18F23J0015 |
| IAB-06830-U | 18F23J0046 | (01)10801902161939(17)250831(10)18F23J0046 |
| IAB-06830-U | 18F23J0053 | (01)10801902172898(17)250831(10)18F23J0053 |
| IAB-06830-U | 18F23J0054 | (01)10801902172898(17)250831(10)18F23J0054 |
| IAB-06830-U | 18F23J0055 | (01)10801902172898(17)250831(10)18F23J0055 |
| IAB-06830-U | 18F23J0079 | (01)10801902172898(17)250930(10)18F23J0079 |
| IAB-06830-U | 18F23K0003 | (01)10801902172898(17)250930(10)18F23K0003 |
| IAB-06830-U | 18F23K0005 | (01)10801902172898(17)250930(10)18F23K0005 |
| IAB-06830-U | 18F23K0009 | (01)10801902172898(17)250930(10)18F23K0009 |
| IAB-06830-U | 18F23K0051 | (01)10801902172898(17)250930(10)18F23K0051 |
| IAB-06830-U | 18F23K0059 | (01)10801902161939(17)250930(10)18F23K0059 |
| IAB-06830-U | 18F23K0061 | (01)10801902172898(17)250930(10)18F23K0061 |
| IAB-06830-U | 18F23L0040 | (01)10801902172898(17)251031(10)18F23L0040 |
| IAB-06830-U | 18F23M0002 | (01)10801902172898(17)251130(10)18F23M0002 |
| IAB-06830-U | 18F23M0044 | (01)10801902172898(17)251130(10)18F23M0044 |
| IAB-06830-U | 18F23M0049 | (01)10801902172898(17)251130(10)18F23M0049 |
| IAB-06830-U | 18F23M0053 | (01)10801902172898(17)251130(10)18F23M0053 |
| IAB-06830-U | 18F23M0054 | (01)10801902172898(17)251130(10)18F23M0054 |
| IAB-06830-U | 18F23M0055 | (01)10801902172898(17)251231(10)18F23M0055 |
| IAB-06830-U | 18F24A0006 | (01)10801902172898(17)251231(10)18F24A0006 |
| IAB-06830-U | 18F24A0012 | (01)10801902172898(17)251231(10)18F24A0012 |
| IAB-06830-U | 18F24A0013 | (01)10801902172898(17)251231(10)18F24A0013 |
| IAB-06830-U | 18F24A0021 | (01)10801902172898(17)251231(10)18F24A0021 |
| IAB-06830-U | 18F24A0036 | (01)10801902161939(17)251231(10)18F24A0036 |
| IAB-06830-U | 18F24B0014 | (01)10801902172898(17)260131(10)18F24B0014 |
| IAB-06830-U | 18F24C0004 | (01)10801902172898(17)260228(10)18F24C0004 |
| IAB-06830-U | 18F24C0005 | (01)10801902172898(17)260228(10)18F24C0005 |
| IAB-06830-U | 18F24D0004 | (01)10801902161939(17)260331(10)18F24D0004 |
| IAB-06830-U | 18F24D0033 | (01)10801902172898(17)260331(10)18F24D0033 |
| IAB-06830-U | 18F24D0036 | (01)10801902172898(17)260331(10)18F24D0036 |
| IAB-06830-U | 18F24D0043 | (01)10801902172898(17)260331(10)18F24D0043 |
| IAB-06830-U | 18F24D0048 | (01)10801902172898(17)260331(10)18F24D0048 |
| IAB-06830-U | 18F24D0086 | (01)10801902172898(17)260331(10)18F24D0086 |
| IAB-06840-U | 18F22E0001 | (01)10801902172904(17)240430(10)18F22E0001 |
| IAB-06840-U | 18F22E0009 | (01)10801902172904(17)240430(10)18F22E0009 |
| IAB-06840-U | 18F22E0014 | (01)10801902172904(17)240430(10)18F22E0014 |
| IAB-06840-U | 18F22E0020 | (01)10801902172904(17)240430(10)18F22E0020 |
| IAB-06840-U | 18F22E0023 | (01)10801902161946(17)240430(10)18F22E0023 |
| IAB-06840-U | 18F22E0036 | (01)10801902172904(17)240430(10)18F22E0036 |
| IAB-06840-U | 18F22E0037 | (01)10801902172904(17)240430(10)18F22E0037 |
| IAB-06840-U | 18F22E0042 | (01)10801902172904(17)240430(10)18F22E0042 |
| IAB-06840-U | 18F22E0046 | (01)10801902172904(17)240430(10)18F22E0046 |
| IAB-06840-U | 18F22E0048 | (01)10801902172904(17)240430(10)18F22E0048 |
| IAB-06840-U | 18F22E0049 | (01)10801902172904(17)240430(10)18F22E0049 |
| IAB-06840-U | 18F22E0061 | (01)10801902172904(17)240430(10)18F22E0061 |
| IAB-06840-U | 18F22E0065 | (01)10801902172904(17)240430(10)18F22E0065 |
| IAB-06840-U | 18F22E0066 | (01)10801902172904(17)240430(10)18F22E0066 |
| IAB-06840-U | 18F22F0002 | (01)10801902172904(17)240531(10)18F22F0002 |
| IAB-06840-U | 18F22F0009 | (01)10801902161946(17)240531(10)18F22F0009 |
| IAB-06840-U | 18F22F0010 | (01)10801902172904(17)240531(10)18F22F0010 |
| IAB-06840-U | 18F22F0016 | (01)10801902172904(17)240531(10)18F22F0016 |
| IAB-06840-U | 18F22F0019 | (01)10801902161946(17)240531(10)18F22F0019 |
| IAB-06840-U | 18F22F0021 | (01)10801902172904(17)240531(10)18F22F0021 |
| IAB-06840-U | 18F22F0022 | (01)10801902172904(17)240531(10)18F22F0022 |
| IAB-06840-U | 18F22F0037 | (01)10801902172904(17)240531(10)18F22F0037 |
| IAB-06840-U | 18F22F0043 | (01)10801902172904(17)240531(10)18F22F0043 |
| IAB-06840-U | 18F22F0047 | (01)10801902172904(17)240531(10)18F22F0047 |
| IAB-06840-U | 18F22F0050 | (01)10801902172904(17)240531(10)18F22F0050 |
| IAB-06840-U | 18F22F0051 | (01)10801902172904(17)240531(10)18F22F0051 |
| IAB-06840-U | 18F22F0054 | (01)10801902172904(17)240531(10)18F22F0054 |
| IAB-06840-U | 18F22F0065 | (01)10801902172904(17)240531(10)18F22F0065 |
| IAB-06840-U | 18F22F0066 | (01)10801902172904(17)240531(10)18F22F0066 |
| IAB-06840-U | 18F22F0067 | (01)10801902172904(17)240531(10)18F22F0067 |
| IAB-06840-U | 18F22G0067 | (01)10801902172904(17)240630(10)18F22G0067 |
| IAB-06840-U | 18F22H0004 | (01)10801902172904(17)240731(10)18F22H0004 |
| IAB-06840-U | 18F22H0010 | (01)10801902172904(17)240731(10)18F22H0010 |
| IAB-06840-U | 18F22H0014 | (01)10801902161946(17)240731(10)18F22H0014 |
| IAB-06840-U | 18F22H0026 | (01)10801902172904(17)240531(10)18F22H0026 |
| IAB-06840-U | 18F22H0028 | (01)10801902161946(17)240731(10)18F22H0028 |
| IAB-06840-U | 18F22H0030 | (01)10801902172904(17)240731(10)18F22H0030 |
| IAB-06840-U | 18F22H0036 | (01)10801902161946(17)240731(10)18F22H0036 |
| IAB-06840-U | 18F22H0038 | (01)10801902172904(17)240731(10)18F22H0038 |
| IAB-06840-U | 18F22H0039 | (01)10801902172904(17)240731(10)18F22H0039 |
| IAB-06840-U | 18F22K0004 | (01)10801902172904(17)240930(10)18F22K0004 |
| IAB-06840-U | 18F22M0015 | (01)10801902161946(17)241130(10)18F22M0015 |
| IAB-06840-U | 18F23A0019 | (01)10801902161946(17)241231(10)18F23A0019 |
| IAB-06840-U | 18F23A0025 | (01)10801902161946(17)250228(10)18F23A0025 |
| IAB-06840-U | 18F23A0037 | (01)10801902161946(17)241231(10)18F23A0037 |
| IAB-06840-U | 18F23A0054 | (01)10801902161946(17)241231(10)18F23A0054 |
| IAB-06840-U | 18F23A0055 | (01)10801902161946(17)241231(10)18F23A0055 |
| IAB-06840-U | 18F23A0061 | (01)10801902161946(17)241231(10)18F23A0061 |
| IAB-06840-U | 18F23A0066 | (01)10801902161946(17)250131(10)18F23A0066 |
| IAB-06840-U | 18F23A0067 | (01)10801902161946(17)250131(10)18F23A0067 |
| IAB-06840-U | 18F23B0018 | (01)10801902161946(17)250228(10)18F23B0018 |
| IAB-06840-U | 18F23C0017 | (01)10801902172904(17)250228(10)18F23C0017 |
| IAB-06840-U | 18F23C0027 | (01)10801902172904(17)250228(10)18F23C0027 |
| IAB-06840-U | 18F23C0044 | (01)10801902172904(17)250228(10)18F23C0044 |
| IAB-06840-U | 18F23D0015 | (01)10801902172904(17)250331(10)18F23D0015 |
| IAB-06840-U | 18F23D0018 | (01)10801902161946(17)250331(10)18F23D0018 |
| IAB-06840-U | 18F23D0027 | (01)10801902172904(17)250430(10)18F23D0027 |
| IAB-06840-U | 18F23D0028 | (01)10801902172904(17)250430(10)18F23D0028 |
| IAB-06840-U | 18F23E0013 | (01)10801902161946(17)250430(10)18F23E0013 |
| IAB-06840-U | 18F23E0014 | (01)10801902172904(17)250430(10)18F23E0014 |
| IAB-06840-U | 18F23E0015 | (01)10801902161946(17)250430(10)18F23E0015 |
| IAB-06840-U | 18F23E0028 | (01)10801902172904(17)250430(10)18F23E0028 |
| IAB-06840-U | 18F23E0035 | (01)10801902172904(17)250430(10)18F23E0035 |
| IAB-06840-U | 18F23E0036 | (01)10801902172904(17)250430(10)18F23E0036 |
| IAB-06840-U | 18F23E0050 | (01)10801902172904(17)250430(10)18F23E0050 |
| IAB-06840-U | 18F23E0051 | (01)10801902172904(17)250430(10)18F23E0051 |
| IAB-06840-U | 18F23E0052 | (01)10801902172904(17)250430(10)18F23E0052 |
| IAB-06840-U | 18F23E0055 | (01)10801902172904(17)250531(10)18F23E0055 |
| IAB-06840-U | 18F23F0003 | (01)10801902172904(17)250531(10)18F23F0003 |
| IAB-06840-U | 18F23F0006 | (01)10801902172904(17)250531(10)18F23F0006 |
| IAB-06840-U | 18F23F0007 | (01)10801902172904(17)250531(10)18F23F0007 |
| IAB-06840-U | 18F23F0014 | (01)10801902172904(17)250531(10)18F23F0014 |
| IAB-06840-U | 18F23F0015 | (01)10801902172904(17)250531(10)18F23F0015 |
| IAB-06840-U | 18F23F0024 | (01)10801902172904(17)250531(10)18F23F0024 |
| IAB-06840-U | 18F23F0025 | (01)10801902172904(17)250531(10)18F23F0025 |
| IAB-06840-U | 18F23F0048 | (01)10801902172904(17)250531(10)18F23F0048 |
| IAB-06840-U | 18F23F0052 | (01)10801902172904(17)250531(10)18F23F0052 |
| IAB-06840-U | 18F23F0054 | (01)10801902172904(17)250531(10)18F23F0054 |
| IAB-06840-U | 18F23F0059 | (01)10801902172904(17)250531(10)18F23F0059 |
| IAB-06840-U | 18F23G0002 | (01)10801902161946(17)250630(10)18F23G0002 |
| IAB-06840-U | 18F23G0003 | (01)10801902161946(17)250630(10)18F23G0003 |
| IAB-06840-U | 18F23G0005 | (01)10801902172904(17)250630(10)18F23G0005 |
| IAB-06840-U | 18F23G0007 | (01)10801902172904(17)250630(10)18F23G0007 |
| IAB-06840-U | 18F23G0011 | (01)10801902161946(17)250630(10)18F23G0011 |
| IAB-06840-U | 18F23G0014 | (01)10801902172904(17)250630(10)18F23G0014 |
| IAB-06840-U | 18F23G0028 | (01)10801902172904(17)250630(10)18F23G0028 |
| IAB-06840-U | 18F23G0050 | (01)10801902172904(17)250630(10)18F23G0050 |
| IAB-06840-U | 18F23G0053 | (01)10801902172904(17)250630(10)18F23G0053 |
| IAB-06840-U | 18F23G0063 | (01)10801902161946(17)250731(10)18F23G0063 |
| IAB-06840-U | 18F23G0064 | (01)10801902172904(17)250731(10)18F23G0064 |
| IAB-06840-U | 18F23G0065 | (01)10801902172904(17)250731(10)18F23G0065 |
| IAB-06840-U | 18F23H0033 | (01)10801902172904(17)250831(10)18F23H0033 |
| IAB-06840-U | 18F23H0034 | (01)10801902161946(17)250731(10)18F23H0034 |
| IAB-06840-U | 18F23H0035 | (01)10801902161946(17)250731(10)18F23H0035 |
| IAB-06840-U | 18F23H0038 | (01)10801902161946(17)250731(10)18F23H0038 |
| IAB-06840-U | 18F23H0053 | (01)10801902161946(17)250731(10)18F23H0053 |
| IAB-06840-U | 18F23H0057 | (01)10801902172904(17)250731(10)18F23H0057 |
| IAB-06840-U | 18F23H0058 | (01)10801902161946(17)250731(10)18F23H0058 |
| IAB-06840-U | 18F23H0071 | (01)10801902161946(17)250731(10)18F23H0071 |
| IAB-06840-U | 18F23H0075 | (01)10801902172904(17)250731(10)18F23H0075 |
| IAB-06840-U | 18F23J0006 | (01)10801902161946(17)250831(10)18F23J0006 |
| IAB-06840-U | 18F23J0008 | (01)10801902161946(17)250831(10)18F23J0008 |
| IAB-06840-U | 18F23J0014 | (01)10801902161946(17)250831(10)18F23J0014 |
| IAB-06840-U | 18F23J0018 | (01)10801902161946(17)250831(10)18F23J0018 |
| IAB-06840-U | 18F23J0019 | (01)10801902172904(17)250831(10)18F23J0019 |
| IAB-06840-U | 18F23J0035 | (01)10801902161946(17)250831(10)18F23J0035 |
| IAB-06840-U | 18F23J0041 | (01)10801902172904(17)250831(10)18F23J0041 |
| IAB-06840-U | 18F23J0042 | (01)10801902161946(17)250831(10)18F23J0042 |
| IAB-06840-U | 18F23J0050 | (01)10801902172904(17)250831(10)18F23J0050 |
| IAB-06840-U | 18F23K0017 | (01)10801902172904(17)250930(10)18F23K0017 |
| IAB-06840-U | 18F23K0018 | (01)10801902172904(17)250930(10)18F23K0018 |
| IAB-06840-U | 18F23K0019 | (01)10801902172904(17)250930(10)18F23K0019 |
| IAB-06840-U | 18F23K0027 | (01)10801902172904(17)250930(10)18F23K0027 |
| IAB-06840-U | 18F23K0028 | (01)10801902161946(17)250930(10)18F23K0028 |
| IAB-06840-U | 18F23K0040 | (01)10801902172904(17)250930(10)18F23K0040 |
| IAB-06840-U | 18F23K0045 | (01)10801902172904(17)250930(10)18F23K0045 |
| IAB-06840-U | 18F23K0057 | (01)10801902172904(17)250930(10)18F23K0057 |
| IAB-06840-U | 18F23K0060 | (01)10801902172904(17)251031(10)18F23K0060 |
| IAB-06840-U | 18F23L0006 | (01)10801902172904(17)251031(10)18F23L0006 |
| IAB-06840-U | 18F23L0029 | (01)10801902172904(17)251031(10)18F23L0029 |
| IAB-06840-U | 18F23L0030 | (01)10801902172904(17)251031(10)18F23L0030 |
| IAB-06840-U | 18F23M0019 | (01)10801902172904(17)251130(10)18F23M0019 |
| IAB-06840-U | 18F23M0020 | (01)10801902172904(17)251130(10)18F23M0020 |
| IAB-06840-U | 18F23M0037 | (01)10801902172904(17)251130(10)18F23M0037 |
| IAB-06840-U | 18F23M0039 | (01)10801902172904(17)251130(10)18F23M0039 |
| IAB-06840-U | 18F23M0040 | (01)10801902172904(17)251130(10)18F23M0040 |
| IAB-06840-U | 18F23M0041 | (01)10801902172904(17)251130(10)18F23M0041 |
| IAB-06840-U | 18F23M0064 | (01)10801902172904(17)251231(10)18F23M0064 |
| IAB-06840-U | 18F23M0065 | (01)10801902172904(17)251231(10)18F23M0065 |
| IAB-06840-U | 18F24A0002 | (01)10801902172904(17)251231(10)18F24A0002 |
| IAB-06840-U | 18F24A0003 | (01)10801902172904(17)251231(10)18F24A0003 |
| IAB-06840-U | 18F24A0004 | (01)10801902172904(17)251231(10)18F24A0004 |
| IAB-06840-U | 18F24A0005 | (01)10801902172904(17)251231(10)18F24A0005 |
| IAB-06840-U | 18F24A0014 | (01)10801902172904(17)251231(10)18F24A0014 |
| IAB-06840-U | 18F24A0016 | (01)10801902172904(17)251231(10)18F24A0016 |
| IAB-06840-U | 18F24A0022 | (01)10801902172904(17)251231(10)18F24A0022 |
| IAB-06840-U | 18F24A0023 | (01)10801902172904(17)251231(10)18F24A0023 |
| IAB-06840-U | 18F24A0024 | (01)10801902172904(17)251231(10)18F24A0024 |
| IAB-06840-U | 18F24A0038 | (01)10801902172904(17)251231(10)18F24A0038 |
| IAB-06840-U | 18F24B0006 | (01)10801902172904(17)260131(10)18F24B0006 |
| IAB-06840-U | 18F24B0007 | (01)10801902172904(17)260131(10)18F24B0007 |
| IAB-06840-U | 18F24B0023 | (01)10801902172904(17)260131(10)18F24B0023 |
| IAB-06840-U | 18F24B0024 | (01)10801902172904(17)260131(10)18F24B0024 |
| IAB-06840-U | 18F24B0025 | (01)10801902161946(17)260131(10)18F24B0025 |
| IAB-06840-U | 18F24B0026 | (01)10801902161946(17)260131(10)18F24B0026 |
| IAB-06840-U | 18F24B0027 | (01)10801902172904(17)260131(10)18F24B0027 |
| IAB-06840-U | 18F24B0028 | (01)10801902172904(17)260131(10)18F24B0028 |
| IAB-06840-U | 18F24B0045 | (01)10801902172904(17)260131(10)18F24B0045 |
| IAB-06840-U | 18F24B0046 | (01)10801902172904(17)260131(10)18F24B0046 |
| IAB-06840-U | 18F24B0047 | (01)10801902172904(17)260131(10)18F24B0047 |
| IAB-06840-U | 18F24B0067 | (01)10801902172904(17)260228(10)18F24B0067 |
| IAB-06840-U | 18F24B0068 | (01)10801902172904(17)260131(10)18F24B0068 |
| IAB-06840-U | 18F24B0069 | (01)10801902172904(17)260131(10)18F24B0069 |
| IAB-06840-U | 18F24B0070 | (01)10801902172904(17)260228(10)18F24B0070 |
| IAB-06840-U | 18F24B0071 | (01)10801902172904(17)260228(10)18F24B0071 |
| IAB-06840-U | 18F24C0006 | (01)10801902172904(17)260228(10)18F24C0006 |
| IAB-06840-U | 18F24C0007 | (01)10801902172904(17)260228(10)18F24C0007 |
| IAB-06840-U | 18F24C0028 | (01)10801902172904(17)260228(10)18F24C0028 |
| IAB-06840-U | 18F24C0029 | (01)10801902172904(17)260228(10)18F24C0029 |
| IAB-06840-U | 18F24C0030 | (01)10801902172904(17)260228(10)18F24C0030 |
| IAB-06840-U | 18F24C0038 | (01)10801902172904(17)260228(10)18F24C0038 |
| IAB-06840-U | 18F24C0092 | (01)10801902172904(17)260228(10)18F24C0092 |
| IAB-06840-U | 18F24C0120 | (01)10801902172904(17)260228(10)18F24C0120 |
| IAB-06840-U | 18F24C0121 | (01)10801902172904(17)260228(10)18F24C0121 |
| IAB-06840-U | 18F24C0141 | (01)10801902172904(17)260228(10)18F24C0141 |
| IAB-06840-U | 18F24C0151 | (01)10801902161946(17)260228(10)18F24C0151 |
| IAB-06840-U | 18F24C0171 | (01)10801902161946(17)260331(10)18F24C0171 |
| IAB-06840-U | 18F24C0177 | (01)10801902161946(17)260331(10)18F24C0177 |
| IAB-06840-U | 18F24C0178 | (01)10801902161946(17)260331(10)18F24C0178 |
| IAB-06840-U | 18F24C0179 | (01)10801902161946(17)260331(10)18F24C0179 |
| IAB-06840-U | 18F24C0180 | (01)10801902161946(17)260331(10)18F24C0180 |
| IAB-06840-U | 18F24C0195 | (01)10801902161946(17)260331(10)18F24C0195 |
| IAB-06840-U | 18F24D0014 | (01)10801902161946(17)260331(10)18F24D0014 |
| IAB-06840-U | 18F24D0017 | (01)10801902161946(17)260331(10)18F24D0017 |
| IAB-06840-U | 18F24D0030 | (01)10801902161946(17)260331(10)18F24D0030 |
| IAB-06840-U | 18R23G0002 | (01)10801902172904(17)250331(10)18R23G0002 |
| IAB-06850-U | 18F22J0020 | (01)10801902144192(17)240831(10)18F22J0020(13)220916 |
| IAB-06850-U | 18F22J0038 | (01)10801902144192(17)240831(10)18F22J0038(13)220928 |
| IAB-06850-U | 18F22J0043 | (01)10801902144192(17)240831(10)18F22J0043(13)220930 |
| IAB-06850-U | 18F22K0001 | (01)10801902144192(17)240930(10)18F22K0001(13)221003 |
| IAB-06850-U | 18F22K0003 | (01)10801902144192(17)240930(10)18F22K0003(13)221004 |
| IAB-06850-U | 18F22K0006 | (01)10801902144192(17)240930(10)18F22K0006(13)221006 |
| IAB-06850-U | 18F22K0011 | (01)10801902144192(17)240930(10)18F22K0011(13)221007 |
| IAB-06850-U | 18F22K0012 | (01)10801902144192(17)240930(10)18F22K0012(13)221007 |
| IAB-06850-U | 18F22K0013 | (01)10801902144192(17)240930(10)18F22K0013(13)221007 |
| IAB-06850-U | 18F23A0075 | (01)10801902144192(17)241231(10)18F23A0075(13)230131 |
| IAB-06850-U | 18F23B0005 | (01)10801902144192(17)250131(10)18F23B0005(13)230203 |
| IAB-06850-U | 18F23B0010 | (01)10801902144192(17)250131(10)18F23B0010(13)230206 |
| IAB-06850-U | 18F23B0029 | (01)10801902144192(17)250131(10)18F23B0029(13)230220 |
| IAB-06850-U | 18F23B0030 | (01)10801902144192(17)250131(10)18F23B0030(13)230220 |
| IAB-06850-U | 18F23C0007 | (01)10801902144192(17)250228(10)18F23C0007(13)230307 |
| IAB-06850-U | 18F23C0008 | (01)10801902144192(17)250228(10)18F23C0008(13)230307 |
| IAB-06850-U | 18F23C0014 | (01)10801902144192(17)250228(10)18F23C0014(13)230310 |
| IAB-06850-U | 18F23C0015 | (01)10801902144192(17)250228(10)18F23C0015(13)230310 |
| IAB-06850-U | 18F23F0016 | (01)10801902144192(17)250531(10)18F23F0016(13)230609 |
| IAB-06850-U | 18F23F0017 | (01)10801902144192(17)250531(10)18F23F0017(13)230609 |
| IAB-06850-U | 18F23H0061 | (01)10801902172911(17)250831(10)18F23H0061(13)230911 |
| IAB-06850-U | 18F23J0043 | (01)10801902144192(17)250831(10)18F23J0043(13)230921 |
| IAB-06850-U | 18F23J0044 | (01)10801902144192(17)250831(10)18F23J0044(13)230923 |
| IAB-06850-U | 18F23J0078 | (01)10801902144192(17)250930(10)18F23J0078(13)231031 |
| IAB-06850-U | 18F23K0015 | (01)10801902144192(17)250930(10)18F23K0015(13)231031 |
| IAB-06850-U | 18F23K0024 | (01)10801902144192(17)250930(10)18F23K0024(13)231014 |
| IAB-06850-U | 18F23L0045 | (01)10801902144192(17)251031(10)18F23L0045(13)231130 |
| IAB-06850-U | 18F23L0052 | (01)10801902144192(17)251130(10)18F23L0052(13)231205 |
| IAB-06850-U | 18F23L0058 | (01)10801902144192(17)251130(10)18F23L0058(13)231205 |
| IAB-06850-U | 18F24B0010 | (01)10801902172911(17)260131(10)18F24B0010(13)240207 |
| IAB-06850-U | 18F24B0011 | (01)10801902144192(17)260131(10)18F24B0011(13)240209 |
| IAB-06850-U | 18F24B0012 | (01)10801902144192(17)260131(10)18F24B0012(13)240209 |

***-End of Document-***